The different things I encountered creating my code and which steps I took accordingly
An automatic ROI detection algorithm

Introduction/Summary
Before going into the details of the different steps in the automated ROI detection, I will first give an overview for better understanding of the mechanisms.

For cryo-ET, we need to create a lamella within the sample at the region of interest (ROI). This ROI is found with a fluorescence light microscope (FLM) which is incorporated in the SEM/FIB. To check whether the ROI is still present after FIB milling, we look at the sample through the FLM and measure the fluorescence at the spot of the ROI. If there is no fluorescence left, we FIB-milled the ROI away, if there is still fluorescence, the FIB-milling was successful. This check is done manually, however it would be nice to automate it. That is the goal of my project, to automate the ROI detection after FIB-milling.

The general workflow of this automated ROI detection goes as follows:
1) As input it uses the z-stack/image before and after milling (roughly around the milling site)
a) Z-stacks get projected with a maximum intensity projection.
b) The images get normalized to be able to compare them to each other.
2) It finds the location of where there has been milled and creates a mask of that location.
a) By looking at the differences in fluorescent intensity before and after milling.
b) Or alternatively by finding the lines of the milling site with line detection.
3) It looks if there is any fluorescent signal left within the mask in the image after milling.
a) Noise is filtered out with binary operations after applying a threshold to the image.
b) Blob detection can be used if your ROI has certain specific features.

Insert picture of overview, maybe one schematic overview and one overview of the real data in that schematic overview (that you know how it looks in real life so to say)

For more information on the cryo-ET and how the FLM is incorporated, see the website of Delmic:
https://www.delmic.com/en/techniques/cryogenic-electron-microscopy

Reading the image or z-stack
needs to include: pictures of histograms and difference max and mean intensity projections?

The dataset consists of z-stacks and single images in the .tiff and .ome.tiff format. When I first started with the dataset, I manually looked at the z-stacks via Odemis viewer and selected a z-slice from the z-stack which was in focus. For each z-stack this was a different slices number so you could not say: from each z-stack, let’s take slice number 10. Then there was the problem of digital outliers, when I first looked at the images through matplotlib, I saw nothing. Later I found out this was because of the outliers. Where the normal values were from 0 to ~500-10.000, the outliers could get values of up to 65.535. Why this exact upper limit? That is because that is exactly 2^(16) which is the highest value a 16-bit image can get. The cutoff for outliers was also manually chosen at first. For some time I just used these manually chosen values to work out the rest of the steps but after some time the automation of these two steps needed to be implemented (partially because the dataset also got bigger).

So for the automation of picking one image from a z-stack, for simplicity I implemented a mean or maximum intensity projection (MIP). This worked nicely but had some drawbacks, mainly on the memory side of the computer. Z-stacks can be a few hundreds MB in size and when going trough some z-stacks, memory garbage pilled up and the code stopped running at some point. Even after implementing a `del all variables` after every iteration, the memory garbage still pilled up. As a solution I implemented two things: a maximum of slices used and garbage collection. So with the first solution, I assumed that the first and last slices of the z-stack are so out of focus they don’t have any useful contribution to the MIP so they can be disregarded. If the z-stack has more slices than the maximum of slices you tell it to use, the most middle maximum number of slices is used. With this solution, I didn’t really needed the second solution anymore for the dataset I had but if some people have really big datasets, the second solution is still useful, so I still implemented it. At the end of the big for loop of all the z-stacks and images, all the temporary data gets deleted and then with the library Garbage Collector Interface (gc) the memory gets cleaned up. At some point in my dataset this caused an improvement of 10% in memory use at full capacity (8GB of RAM). Another drawback of MIP was that with maximum intensity projection, you have more outlier noise in your image and with mean intensity projection the signal from the in focus slice gets “pulled down” by the absence of signal in the out of focus slices. After looking at both options, I concluded that the maximum intensity projection was better than the mean intensity projection because outlier noise was more easy to filter out than dealing with the lower signal-to-noise ratio (see Figure 1).

Figure 1: Of the same z-stack after milling on the left a mean intensity projection and on the right a max intensity projection. The threshold at the mean intensity projection had to be set at 0.075 to see the blob in the middle and th threshold at the max intensity projection was set to 0.25. Here you can clearly see that you can more easily detect signal in a max intensity projection.

For the automation of the outlier threshold, I first looked if I could just use the 99% of the maximum intensity as a threshold. However this did not work for all the z-stacks as the maximum values varied quite a lot. As a second approach, I tried to look at the histogram of the intensity values. I made a cumulative normalized histogram and made a cutoff at 99% of all the pixels. This worked way better and often even gave the same threshold as was set manually. So this was in the end the way to go for the outlier cutoff threshold. See figure 2 below for the cumulative histograms before and after the cutoff, pay attention to the x-axis.

Figure 2: The cumulative histogram before and after the cutoff has been used on the image. You can see that the cutoff was somewhere around 700 and that the pixels above that value are only a really small portion of the entire image (that really small peak at the end).
Re-scaling and shifting
Already very soon I noticed that sometimes the z-stack before milling had a different resolution and image pixel size than the z-stack after milling. To be able to look at the differences between the two images, they need the same pixel size. As a solution, I implemented an interpolation such that the lower resolution image had the same number of pixels as the higher resolution image. In this way, nothing qualitatively changed and the images could be properly compared.

Next to the different resolutions / pixel sizes, there was also a shift in the image after milling compared to the image before milling. This is due to the fact that the stage needs to move towards the SEM/FIB place and back towards the METEOR spot. The mechanical uncertainties give rise to a shift in the image after the milling. So, this needs to be accounted for, but how? First, I wanted to use the metadata of the images to calculate the shift, but after some experimentation I noticed that the metadata was mostly not correct and not accurate (especially in the z-axis position, there were shifts found of hundreds slices while the z-stacks only had 30-70 slices, see figure 3). So instead of using the metadata, I implemented a convolution between the two images. With finding the maximum in the convolution, you could find the shift of the image after milling. This worked nicely and seemed to be more accurate. However for some images, due to the shift, other fluorescent signal came into the image, which sometimes gave rise to the wrong shift. To account for this, I assumed that the images roughly overlap and that shifts bigger than 1/x of the image cannot occur (so the best next option gets chosen). This x is a parameter that you can change yourself (in the function itself the 1/x works a bit different but the essence is the same). If in the future with new data, this convolution method does not work, SIFT could be a good alternative (not implemented).

After the image after milling was shifted to the right position, then the pixels that fell outside of the shifted image boundary were given the maximum intensity value of the image. In this way, these pixels would not be used anymore in the next steps of the automated ROI detection.
Figure 3: Different shifts from different images calculated via the convolution and via the metadata.

Normalizing and blurring
After milling, the fluorescence in the sample at that spot has decreased significantly. Because of this, researchers sometimes use a higher illumination intensity to still see the fluorescence above the noise. This works for the researchers but here if you want to compare the image before and after milling, the intensity scales are sometimes very different (see figure 4). So we need to account for these different intensity scales. This is done with a normalization in both images to go to a range of 0 to 1. In this way we can easily compare them and look at the differences quantitatively (see figure 5). I actually did not use the full range of the images to normalize it. This is because there is a constant background intensity which we can get rid of with finding the minimum intensity of a blurred image. Why the blurred image? Because with noise you still have relatively large variations in this background intensity (e.g. you have a background level of 70 but with the noise the minimal intensity is 0). With blurring the image, you account for the noise and get rid of it. Then you do find the minimum to be the background intensity level instead of just zero. In this way, going from background level to max intensity instead of going from zero to max intensity, the highest contrast is achieved with the normalization. Everything below the background level is just set to zero because that is probably noise.

Figure 4: Intensity histogram before and after milling without normalization. The image after milling has higher intensity values than the image before milling.

Figure 5: The normalized intensity histogram of the images before and after milling. The intensity profiles are quite similar, which we want to compare the two.

Next to the normalization, the blurred images are also used for detecting where the ROI is with the differences in intensity. This is because the location detection is not detail orientated so blurred images are preferred (then the noise is better suppressed). See next section for more information.

Mask creation with the difference
So the first way to create a mask of the ROI location is to look at the differences between the image before and after milling. Here I used the blurred images because you want to look at larger regions of intensity differences and not at very small spots. With blurring you get rid of small local intensity changes (mostly caused by noise). The difference between the images is calculated with just a simple subtraction of before minus after. A threshold is applied to the difference and we get a boolean image which represents the mask. Sometimes, small spots of intensity differences still get through the threshold so a few binary operations are performed on the mask to get rid of too small regions. First, a binary opening is performed to get rid of too small regions in the boolean mask (see figure 6). Secondly, an binary erosion can be optionally performed to get rid of signal on the boundaries of the mask. Lastly, the mask can be made a square to represent the milling site more closely, this is done with finding the maxima and minima of the mask and use that as the boundaries of the square (see figure 6). In most cases it seems that squaring the mask is beneficial, because if the intensity after milling is still too high, the difference is too low and it will not be taken into the mask (see figure 7). However, sometimes there are still multiple regions in the mask that create a square almost as big as the image itself (see figure 8). As a solution to this, I implemented a constraint to how big the mask can be so if squaring makes it bigger than the maximum, it will not be squared.

Figure 6: From left to right, the differences in intensity; the mask after thresholding without any binary operations; the mask with a binary opening performed; the mask made into a square. Here you can see that before the binary opening there still is a small spot at the bottom of the image which is gone after the binary opening.

Figure 7: The masks created overlapping with the images after milling. Here you can see that where the image has too much intensity left, the mask will form around that spot instead of taking it with it.
Figure 8: An image with multiple regions in the mask which will give a wrong mask if squared.

Line detection for mask creation
Because the mask created with the intensity difference is not always optimal, I looked at another way of finding the milling site: line detection. Because of the milling, two large lines are seen in the image. With finding these lines, we can find the milling site. I implemented a line detection with the opencv Hough lines detection (https://docs.opencv.org/3.4/d3/de6/tutorial_js_houghlines.html). I first tried to find these two big lines in one go, however the biological samples were not straight enough in their lines to detect it. There were too much gaps and curves. When letting those curves and gaps be tolerated, too much lines were detected throughout the whole image (see figure 9). My second approach was to detect smaller lines and group them together. First, because the lines are small, there are still a lot of lines to work with so a preselection is done on if it is roughly around the estimated x-location of the milling site (see figure 10). Then the lines are combined into one if they are too close to each other and have the same angle (see figure 11). After that the lines are grouped together on distance from each other when you extend each line to the full image and in this way you could see it as pasting the smaller line pieces together into one big line (if possible). With the bigger lines (consisting of the small lines), I looked at which of those could form a milling site: having a distance of roughly the width of a milling site from each other. From the bigger line pairs, the one pair consisting of the most smaller lines is chosen to be the final answer (see figure 12). If there are multiple pairs with the same most number of small lines, first it is tried to combine the pairs and look again for the most small lines. If that doesn’t work, the one closest to the estimated x-position of the milling site is chosen. In this way, the milling site was estimated correctly the most. For creating the mask, all pixels between these two lines are chosen to be the mask.

So in short the line detection goes as follows:
1) Find small lines
2) Group small lines into bigger lines
3) Pair the bigger lines
4) Choose the pair with the most small lines
5) Create the mask

However, still some estimations were incorrect with off-sites (see figure 13). This is because for example in yeast samples, the milling lines are not well seen and the number of small lines is very low. It is difficult to fine tune all the different parameters in the line detection to have an optimal detection, so that is something I will try again later in the development as finalization.

Figure 9: The line detection in an image after milling with directly searching for big lines. It gave a lot of lines and not even at the right place (on the right side of the milling site).

Figure 10: On the left the line detection without any processing done to it and on the right the preselection of lines around the estimated x-position of the milling site and the combination of similar lines.

Figure 11: On the left small lines without any combination of the lines done and on the right you can see that the similar lines are combined into one.

Figure 12: On the left the lines before grouping and final selection and on the right after grouping and selecting for the pair of groups with the biggest number of smaller lines.

Figure 13: Here you can an off-site with the line detection on the right. The milling site is on the right of that clump of cells.
Mask creation with the lines and the difference
So now we have one mask created with the differences in intensity and one mask created with line detection. But which one is better to use? Because the line detection mask can sometimes be at a different place entirely, this is the lesser one of the two. However, it would be still nice to combine the two because they have both their strengths and weaknesses:

Because both have very different approaches to locating the milling site, if the estimations do overlap, this is probably the correct estimation. So I created a function that tries to combine the two masks. If there is enough overlap of the two masks, then the overlap is taken to be the new mask and dilated in the x-direction to the line detection mask. It is also dilated in the y-direction but not too much because otherwise there would be unnecessary signal of off-sites within the mask (above and below the milling site). If there is not enough overlap, only the mask of the intensity difference is used. See figure 14 for one example of the masks of a sample in which the combination is done and figure 15 for one example of where the combination is not done.

Figure 14: From left to right: The mask of the intensity difference, the mask of the line detection, the overlap of the two masks and the masks combined. Here you see that the two masks overlap so they are combined.

Figure 15: From left to right: The mask of the intensity difference, the mask of the line detection, the overlap of the two masks and the masks combined. Here you see that there is no overlap so only the mask of the intensity difference is used.

Final threshold for detecting signal within the mask
So now we have the location of the milling site within the mask. The next step is to search for signal within this mask in the image after milling. This is done with setting a simple threshold. However, then still some noise gets through the threshold so some binary operations have to be performed. First, to get rid of the noise, a binary opening is performed. However, it was seen that if a big blob had some small holes in it, this opening would destroy the blob, so first a small binary closing was done to close holes in bigger objects (see figure 16).

Figure 16: Top Left) No binary operations are performed. Top Right) Only a binary opening is performed. Bottom Left) Only a binary closing is performed. Bottom Right) First a closing and then an opening is performed. Here we can see that it is beneficial to do first the closing and then the opening.

So now we have first a small closing and then a bigger opening. After this, it was seen that at the boundaries of the mask, we had signal from cells outside the milling site. To group those signals together, another binary closing was performed and then you could get rid of it all together with selecting for signal at the boundaries. See figure 17 below for the different steps in this procedure.

Figure 17: in three different samples from left to right: no binary operations; only the first binary closing; first binary closing, then binary opening; first binary closing, then binary opening and as last again a binary closing.

After these binary operations, you can choose to get rid of signal on the boundary or not. I implemented an option in the function which does these binary operations that you can get rid of the boundary signal. You want to do this because this signal is probably from just outside of the milling site. This is not the signal you are looking for. See figure 18 below for the differences.

Figure 18: On the left the binary end result with the boundary signal and on the right without the boundary signal.

Circle detection became blob detection
Actually before getting rid of the boundary signal, I had a different way of excluding that signal: blob detection from opencv (https://learnopencv.com/blob-detection-using-opencv-python-c/). I would select for the roundness of blobs and in that way you could get rid of the boundary signal, which is in general very elongated as seen in figure 18. I also implemented that the blob needs a minimum area to be detected. In this way it makes it more robust that you are more certain about the fact that it is actual signal. For now these are the only two constraints added but I made it in such a way that you can fine tune the selection on the features your ROI has. You can also implement a maximum area, a minimum and maximum inertia and a minimum and maximum convexity. Inertia is how elongated a blob is and convexity is defined as the (area of the blob / area of it’s convex hull). The convex hull of a shape is the tightest convex shape that completely encloses the shape. All have a range from 0 to 1 except the area, that is from 0 to infinity. There is one other feature which we can select on and that is setting a threshold, however we work here with binary images so no threshold needs to be set. See figure 19 below for examples of the different features and see figure 20 for how the blob detection works on the result from the final threshold. As you can see the right image of figure 20, holes in blobs and small roundish signal at the boundary are also detected while those shouldn’t be detected. Because of this, the method of getting rid of the boundary signal in total as shown before is more beneficial. After that, still a selection on features can be done with the detect_blobs function but it depends on what you look for in a sample.

Figure 19: The different features you can select on with the blob detection from opencv.

Figure 20: The white images are the result from the final threshold discussed in the last section, the red circles are the blobs detected with the specified constraints of: minimal circularity of 0.6 and minimal area of 25 pixels.

Plotting end results
In the end I make an overview figure with the images at different steps in the progress. In these images, on the left you see the input images before and after milling which comes from the z-projection or the found in focus z-slice. On the right of that, you see the blurred after milling image with a rectangle at the found milling site of the intensity difference mask and the combined mask. Then on the right of that, you see the image after milling within these masks. And last on the right, you see the binary end result after the threshold and binary operations were applied. See figure 21 for an example of such an overview image.

Figure 21: From left to right: the input images before and after milling; the blurred after milling image with a rectangle at the found milling site of the combined and the intensity difference mask; The image after milling within the combined and the intensity difference mask; the end result from applying a threshold and binary operations.